## Data Qualifiers Table
Corrective Action and DMR Qualification Table for WWTP Basic Four Tests [5.5.11]
Test | Corrective Action Must Be Taken and Documented | DMR Must Be Qualified | Example of How Best to qualify this data on the DMR
For All Tests | For All Tests | For All Tests | For All Tests
Your sample acceptance policy must indicate under what conditions samples received will require the data be qualified. If these conditions exist then | Yes contact client | Yes | Results [indicate test affected] from sample date mm/dd/yy were received in "state condition"
Samples received with improper chemical or temperature preservation, beyond holding time, inappropriate containers, insufficient volume | Yes contact client | Yes | Results [indicate test affected] from sample date mm/dd/yy were received in "state condition"
QCS (Quarterly Blind Standard) failed | Yes | No | NA
PT standard failed | Yes | No | NA
IDC failed | Yes | No | NA
LOD study failed (not applicable to TSS or BOD) | Yes | No | NA
TSS | TSS | TSS | TSS
Analytical balance - weight verification failed | Yes | No | NA
Insufficient residue captured (< 1.0 mg) | Yes | No | NA
Too much residue captured (> 200 mg) | Yes | No | NA
Quarterly Re-dry study failed | Yes | No | NA
BOD | BOD | BOD | BOD
DO meter calibration failed | Yes | Yes | Samples should never be analyzed with a failing DO meter calibration. In the rare rare event that this must happen the data should be qualified indicating what sample dates were affected
Blank depleted 0.25 mg/L or more | Yes | Yes | The BOD blank associated with sample dates mm/dd/yy [list all dates that apply] failed.
GGA failed | Yes | Yes | The BOD GGA std associated with sample dates mm/dd/yy [list all dates that apply] failed high or low.
None of the seed control dilutions passed | Yes | No | NA
None of the BOD dilutions passed | Yes | No | NA
Samples were supersaturated (> 0.4 TDO or > 9.0 mg/L) | Yes | No | NA
NH3-N | NH3-N | NH3-N | NH3-N
Initial calibration failed | Yes | Yes | Samples should never be analyzed using a failed initial calibration. The calibration issue is to be fixed and samples reanalyzed. In the rare rare event that this must happen the data should be qualified indicating what sample dates were affected by the failed calibration for NH3-N
CCV failed (use when more than 20 samples analyzed in a batch) | Yes | Yes | Samples should never be reported using a failed CCV. The CCV issue is to be fixed and samples reanalyzed. In the rare rare event that this must happen the data must be qualified indicating what sample dates were affected by the failed CCV for NH3-N
Method blank failed | Yes | Yes | NH3-N sample dates mm/dd/yy [list all dates that apply] were reported with a failed method blank
LCS failed | Yes | Yes | NH3-N sample dates mm/dd/yy [list all dates that apply] were reported with a failed LCS
Results reported when the response was higher than the calibration range | Yes | Yes | NH3-N sample dates mm/dd/yy [list all dates that apply] were reported with a result that was not within the calibration range
TP | TP | TP | TP
Initial calibration failed | Yes | Yes | Samples should never be analyzed using a failed initial calibration. The calibration issue is to be fixed and samples reanalyzed. In the rare rare event that this must happen the data should be qualified indicating what sample dates were affected by the failed calibration for TP
ICV failed | Yes | Yes | Samples should never be reported using a failed ICV. The ICV issue is to be fixed and samples reanalyzed. In the rare rare event that this must happen the data must be qualified indicating what sample dates were affected by the failed ICV for TP
CCV failed | Yes | Yes | Samples should never be reported using a failed CCV. The CCV issue is to be fixed and samples reanalyzed. In the rare rare event that this must happen the data must be qualified indicating what sample dates were affected by the failed CCV for TP
Method blank failed | Yes | Yes | TP sample dates mm/dd/yy [list all dates that apply] were reported with a failed method blank
LCS failed | Yes | Yes | TP sample dates mm/dd/yy [list all dates that apply] were reported with a failed LCS
Results reported when the response was higher than the calibration range | Yes | Yes | TP sample dates mm/dd/yy [list all dates that apply] were reported with a result that was not within the calibration range