## Estimate Initial LOD
Estimating Initial LOD Worksheet (Single Instrument) [Version 6/18/24]
Use this form if: |  | a new method is implemented
Use any of the six options to estimate the initial LOD. Fill in the white cells. Cells in yellow have the estimated LOD.
Previously Determined LOD: |  | Instrument limitations or vendor claims: |  | Point at which there is a significant change in sensitivity of the calibration curve: |  | Concentration where Signal to Noise Ratio is 3 - 5:
Spiked Blanks |  |  |  | Method Blanks
| Date Analyzed | Result |  |  | Date Analyzed | Result
1 |  |  |  | 1
2 |  |  |  | 2
3 |  |  |  | 3
4 |  |  |  | 4
5 |  |  |  | 5
6 |  |  |  | 6
7 |  |  |  | 7
8 |  |  |  | 8
Average:
| Standard Deviation: |  |  | Standard Deviation:
| Estimated LODs: |  | mg/L | Estimated LODb: |  |  | mg/L

## Initial LOD (8 MB)
Initial LOD/LOQ Calculation and Validation Worksheet (Single Instrument) [Version 6/18/24]
Use this form if: |  | a new method is implemented.
a significant change has occurred and an initial LOD must be performed again.
there are not enough data to perform the Ongoing Annual Verification
Analytical Method: |  |  |  |  | Calculation Date:
Matrix: |  |  |  |  | Calculation Analyst:
Spiked Blanks | [5 empty] | Method Blanks
| Date Prepped | Date Analyzed | Result | % Recovery |  |  | Date Prepped | Date Analyzed | Result
1 | [5 empty] | 1
2 | [5 empty] | 2
3 | [5 empty] | 3
4 | [5 empty] | 4
5 | [5 empty] | 5 |  |  |  | 0
6 | [5 empty] | 6 |  |  |  | All text
7 | [5 empty] | 7 |  |  |  | 0
8 | [5 empty] | 8
0
Spike Level: |  |  |  | mg/L
Average: | [5 empty] | Average: |  |  | NA | 0
Standard Deviation: |  |  |  |  | 0 | Standard Deviation: |  |  | NA | 0
Student's t-value to use: |  |  |  | 0 | 0 | Student's t-value to use: |  |  | NA | 0
Calculated LODs: |  |  | <7 Spk Blks | mg/L |  | Calculated LODb: |  |  | <7 MBs | mg/L
0 | 0 |  | Calculated LOD: |  | NA | mg/L | (calculated from the greater of LODs and LODb)
|  |  | Calculated LOQ*: |  | NA | mg/L | ( = 10/3 x LOD)
Were outliers rejected? |  |  | (only gross failures may be excluded) | [5 empty] | Yes | No
If so, explain:
NOTES:
The lab may use most recent spiked blank data if generated within the last 2 years and used the same spike level (include initial results if within 2 years).
If the lab thinks the sensitivity of the method has changed significantly, then the most recent data may be used (min. of 7 reps, 3 batches, over 3 days).
Only use data associated with passing calibrations and passing batch QC (reported data).
The LODs and LODb will need to be recalculated at least every 13 months.
*The 2019 NR 149 code requires the LOQ be 10/3 the LOD or at the concentration of the lowest standard in the initial calibration.
LOD Checks (using outdated LOD regulations, but may give insight to LOD determination)
OPTIONAL CHECK Low Spiked Blank Check - Did you spike too low? - - LOD < Spike Concentration - |  |  | LOD | Spike Conc. |  |  |  | If fails, you should spike at higher a concentration | Variability is too high, so the standard deviation is too close to LOD
|  |  | <7 Spk Blks | 0
OPTIONAL CHECK High Spiked Blank Check - Did you spike too high? - - Spike Concentration < 10 x LOD - |  |  | Spike Conc. | LOD x 10 |  |  |  | If fails, you should spike at a lower concentration | LOD is unrealistically low
|  |  | 0 | NA
Regulatory Limit Check - LOD < Permit Limit? - |  |  | LOD | Permit Limit | Enter regulatory limit or permit limit |  |  | If fails, the method must be optimized to be able to meet permit or regulatory limits
NA
OPTIONAL CHECK Is average recovery reasonable? |  |  | Lower Limit | Upper Limit |  |  |  | General guidelines, don't need to re-run study if outside limits
|  |  | 80% | 120%

## Initial LOD (>8 MB)
Initial LOD/LOQ Calculation and Validation Worksheet (Single Instrument) [Version 6/18/24] | [10 empty] | Method Blank Data |  | (this spreadsheet can calculate up to 200 method blanks)
Use this form if: |  | a new method is implemented. | [8 empty] | Date | Result
a significant change has occurred and an initial LOD must be performed again.
there are not enough data to perform the Ongoing Annual Verification
Analytical Method: |  |  |  |  | Calculation Date:
Matrix: |  |  |  |  | Calculation by Analyst:
| 0 | 0
Spiked Blanks | [5 empty] | Method Blanks (see data to the right)
| Date Prepped | Date Analyzed | Result | % Recovery |  | >100 numeric method blank results? |  |  |  | No
1 | [5 empty] | If yes, 99th percentile of blank results: |  |  |  | NA
2
3 | [5 empty] | Average of all results: |  |  |  | NA
4 | [5 empty] | Standard Deviation of all results: |  |  |  | NA
5 | [5 empty] | 0 | All text | 0 |  | Standard Deviation
6
7
8
[6 empty] | LODb if 99th percentile is used: |  |  | NA | mg/L
[6 empty] | LODb if standard deviation is used: |  |  | NA | mg/L
Spike Level: |  |  |  | mg/L |  | Which option will be used: |  |  | Standard Deviation
Average: |  |  | #DIV/0! | #DIV/0! | [5 empty] | (can only use 99th Percentile if more than 100 method blanks)
Standard Deviation: |  |  | #DIV/0! |  |  | 0
Student's t-value to use: |  |  | #NUM! | 0 |  | Student's t-value to use: |  |  | #NUM! | 0
Calculated LODs: |  |  | <7 Spk Blks | mg/L |  | Calculated LODb: |  |  | <7 MBs | mg/L
0
0 | 0 |  | Calculated LOD: |  | NA | mg/L | (calculated from the greater of LODs and LODb)
|  |  | Calculated LOQ*: |  | NA | mg/L | ( = 10/3 x LOD)
Were outliers rejected? |  |  | (only gross failures may be excluded) | [5 empty] | Yes | No
If so, explain:
NOTES:
May use most recent spiked blank data if generated within the last 2 years and used the same spike level (include initial results if within 2 years).
If the lab thinks the sensitivity of the method has changed significantly, then the most recent data may be used (min. of 7 reps, 3 batches, over 3 days).
The lab has the option to use only the last 6 months or the 50 most recent method blanks, whichever yeilds the greater number of blanks.
Only use data associated with passing calibrations and passing batch QC (reported data).
The LODs and LODb will need to be recalculated at least every 13 months.
*The 2019 NR 149 code requires the LOQ be 10/3 the LOD or at the concentration of the lowest standard in the initial calibration.

## Ongoing LOD
Ongoing LOD/LOQ Calculation and Verification Worksheet (Single Instrument) [Version 6/18/24] | [10 empty] | Method Blank Data |  | (this spreadsheet can calculate up to 200 method blanks)
Analytical Method: |  |  |  |  | Calculation Date: | [5 empty] | Date | Result
Matrix: |  |  |  |  | Calculation by Analyst:
| 0 | 0
Spiked Blanks (include data generated within the last 2 years) | [5 empty] | Method Blanks (see data to the right)
| Date Prepped | Date Analyzed | Result | % Recovery |  | >100 numeric method blank results? |  |  |  | No
Initial LOD 1 | [5 empty] | If yes, 99th percentile of blank results: |  |  |  | NA
Initial LOD 2
Initial LOD 3 | [5 empty] | Average of all results: |  |  |  | NA
Initial LOD 4 | [5 empty] | Standard Deviation of all results: |  |  |  | NA
Initial LOD 5 | [5 empty] | All text |  |  |  | 0
Initial LOD 6 | [5 empty] | 0 | Number of text
Initial LOD 7 | [5 empty] | 0 | Number of numeric values |  |  | Standard Deviation
Initial LOD 8 | [5 empty] | 0 | Total MBs
Q1 A
Q1 B
Q2 A
Q2 B
Q3 A
Q3 B
Q4 A
Q4 B
Q1 A
Q1 B
Q2 A
Q2 B
Q3 A
Q3 B
Q4 A | [5 empty] | LODb if 99th percentile is used: |  |  | NA | mg/L
Q4 B | [5 empty] | LODb if standard deviation is used: |  |  | NA | mg/L
Spike Level (same as initial): |  |  |  | mg/L |  | Which option will be used: |  |  | Standard Deviation
Average: |  |  | #DIV/0! | #DIV/0! | [5 empty] | (can only use 99th Percentile if more than 100 method blanks)
Standard Deviation: |  |  | #DIV/0!
Student's t-value to use: |  |  | #NUM! | 0 |  | Student's t-value to use: |  |  | #NUM! | 0
Calculated LODs: |  |  | <7 Spk Blks | mg/L |  | Calculated LODb: |  |  | <7 MBs | mg/L
0
| 0 | 0 | Calculated LOD: |  | NA | mg/L | (calculated from the greater of LODs and LODb)
|  |  | Calculated LOQ: |  | NA | mg/L | ( = 10/3 x LOD)
[7 empty] | 0 | 0 | NO
0 | 0 | #DIV/0! | Existing LOD: |  |  | mg/L | 0 | 0 | 0% | YES
| Can the existing LOD be left unchanged? |  |  |  | Enter existing LOD
|  |  |  | Is spike level okay (<5% spiked blanks <0)? | #DIV/0! |  | #DIV/0!
Were outliers rejected? |  |  | (only gross failures may be excluded) | [5 empty] | Yes | No
If so, explain:
NOTES:
Recalculate LODs and LODb at least every 13 months.
Include spiked blank data generated within the last 2 years as long as the data all used the same spike level (include initial results ONLY if within 2 years). ALL LINES DO NOT HAVE TO BE FILLED IN.
If the lab thinks the sensitivity of the method has changed significantly, then the most recent data may be used (min. of 7 reps, 3 batches, over 3 days).
Include method blank data generated within the last 2 years. The lab has the option to use only the last 6 months or the 50 most recent method blanks, whichever yeilds the greater number of blanks.
The existing LOD can be left unchanged if the calculated LOD is within 0.5 - 2 times the existing LOD AND less than 3% of the method blanks are greater than the existing LOD.
If sample analysis is not performed often for a given test, only quarters with actual sample analysis (do not include PT samples) need to have the spiked blanks analyzed. Take the last 24 months of data to generate the LOD.
If the method is altered in a way that could be expected to change its sensitivity, then re-determine the initial LOD and restart the ongoing data collection.
Only use data associated with passing calibrations and passing batch QC (reported data).
*The 2019 NR 149 code requires the LOQ be 10/3 the LOD or at the concentration of the lowest standard in the initial calibration.

## Add or Change Instrument
Add/Change Instrument LOD/LOQ Calculation and Verification Worksheet [Version 6/18/24]
Use this form if: |  | adding an instrument
Analytical Method: |  |  |  |  | Calculation Date:
Matrix: |  |  |  |  | Calculation by Analyst:
Spiked Blanks (include data generated within the last 2 years) | [5 empty] | Method Blanks
| Date Prepped | Date Analyzed | Result | % Recovery |  |  | Date Prepped | Date Analyzed | Result
Q1 A | [5 empty] | New Inst 1
Q1 B | [5 empty] | New Inst 2
Q2 A
Q2 B
Q3 A
Q3 B
Q4 A
Q4 B
Q1 A
Q1 B | [7 empty] | MB count: | 0
Q2 A | [8 empty] | All text
Q2 B
Q3 A
Q3 B
Q4 A
Q4 B
New Inst 1
New Inst 2
Spike Level (same as initial): |  |  |  | mg/L
Average: |  |  | #DIV/0!
Standard Deviation: |  |  | #DIV/0!
Student's t-value to use: |  |  | #DIV/0! | 0
Calculated LODs: |  |  | #DIV/0! | mg/L
|  |  | Existing LOD: |  |  | mg/L | 0 | 0
|  |  | Is the LODs Valid? |  | Enter existing LOD
|  |  | Is the LODb Valid? |  | Enter existing LOD
|  |  | Is the existing LOD Valid? |  | Enter existing LOD |  | Enter missing data
NOTES:
If a new instrument is added, analyze at least 2 spiked blanks and at least two method blanks on the new instrument.
If the new LODs is within 0.5-2 times the existing LOD, then the LODs is validated.
If the method blanks are less than the existing LOD, then the LODb is validated.
Repeat the initial LOD if either LODs or LODb is not validated.
This form will help determine if the method sensitivity (e.g. LOD) is affected.

## NOTES
Ongoing LOD/LOQ Calculation and Verification Worksheet (Single Instrument) [Version 6/18/24] | [10 empty] | Yes | Method Blank Data |  | (this spreadsheet can calculate up to 200 method blanks)
Analytical Method: |  |  |  |  | Calculation Date: |  |  | 2018-01-22 |  |  | No | Date | Result
Matrix: |  | Wastewater |  |  | Calculation by Analyst: |  |  | AF |  |  | 99th Percentile | 2018-01-02 | 0.0012
| 16 | 16 | [8 empty] | Standard Deviation | 2018-01-02 | 0.0011
Spiked Blanks (include data generated within the last 2 years) | [5 empty] | Method Blanks (see data to the right) | [5 empty] | 2018-01-04 | 0.0054
| Date Prepped | Date Analyzed | Result | % Recovery |  | >100 method blank results? |  |  |  | Yes |  | 2018-01-04 | 0.001
Initial LOD 1 | [5 empty] | If yes, 99th percentile of blank results: |  |  |  | 0.3416 |  | 2018-01-05 | 0.0015
Initial LOD 2 | [11 empty] | 2018-01-05 | 0.0015
Initial LOD 3 | [5 empty] | Average of all results: |  |  |  | 0.0068915 |  | 2018-01-05 | -0.001
Initial LOD 4 | [5 empty] | Standard Deviation of all results: |  |  |  | 0.043466 |  | 2018-01-06 | -0.0072
Initial LOD 5 | [5 empty] | At least some numbers | [5 empty] | 2018-01-07 | 0.001
Initial LOD 6 | [5 empty] | 0 | [5 empty] | 2018-01-08 | 0.0017
Initial LOD 7 | [5 empty] | 132 | [5 empty] | 2018-01-09 | 0.0007
Initial LOD 8 | [11 empty] | 2018-01-10 | 0.00998
Q1 A | 2018-01-04 | 2018-01-04 | 1 | 100% | [7 empty] | 2018-01-11 | 0.00128
Q1 B | 2018-01-11 | 2018-01-11 | 1 | 100% | [7 empty] | 2018-01-12 | 0
Q2 A | 2018-04-05 | 2018-04-05 | 0.888 | 88.8% | [7 empty] | 2018-01-13 | 0.00221
Q2 B | 2018-04-22 | 2018-04-22 | 1.004 | 100.4% | [7 empty] | 2018-01-14 | 0
Q3 A | 2018-07-06 | 2018-07-06 | 0.895 | 89.5% | [7 empty] | 2018-01-15 | 0.0001
Q3 B | 2018-07-13 | 2018-07-13 | 0.904 | 90.4% | [7 empty] | 2018-01-16 | 0.0048
Q4 A | 2018-10-11 | 2018-10-11 | 1.005 | 100.5% | [7 empty] | 2018-01-17 | 0.0067
Q4 B | 2018-10-18 | 2018-10-19 | 1.05 | 105% | [7 empty] | 2018-01-18 | 0.00181
Q1 A | 2019-01-06 | 2019-01-06 | 0.991 | 99.1% | [7 empty] | 2018-01-19 | 0.01
Q1 B | 2019-01-07 | 2019-01-07 | 0.997 | 99.7% | [7 empty] | 2018-01-20 | 0.5
Q2 A | 2019-04-02 | 2019-04-02 | 0.877 | 87.7% | [7 empty] | 2018-01-21 | 0.0012
Q2 B | 2019-04-08 | 2019-04-08 | 0.951 | 95.1% | [7 empty] | 2018-01-22 | 0.0011
Q3 A | 2019-07-04 | 2019-07-04 | 0.967 | 96.7% | [7 empty] | 2018-01-23 | 0.0054
Q3 B | 2019-07-11 | 2019-07-11 | 1.002 | 100.2% | [7 empty] | 2018-01-24 | 0.001
Q4 A | 2019-10-02 | 2019-10-02 | 1.006 | 100.6% |  | LODb if 99th percentile is used: |  |  | 0.3416 | mg/L |  | 2018-01-25 | 0.0015
Q4 B | 2019-10-05 | 2019-10-05 | 0.995 | 99.5% |  | LODb if standard deviation is used: |  |  | 0.10925 | mg/L |  | 2018-01-26 | 0.0015
[12 empty] | 2018-01-27 | 0.02
Spike Level (same as initial): |  |  | 1 | mg/L |  | Which option will be used: |  |  | 99th Percentile |  |  | 2018-01-28 | 0.001
Average: |  |  | 0.97075 | 97.08% | [5 empty] | (can only use 99th Percentile if more than 100 method blanks) |  | 2018-01-29 | 0.001
Standard Deviation: |  |  | 0.051904 | [8 empty] | 2018-01-30 | 0.0017
Student's t-value to use: |  |  | 2.602 | 16 |  | Student's t-value to use: |  |  | 2.355 | 132 |  | 2018-01-31 | 0.0007
Calculated LODs: |  |  | 0.13506 | mg/L |  | Calculated LODb: |  |  | 0.3416 | mg/L |  | 2018-02-01 | 0.00998
[12 empty] | 2018-02-02 | 0.00128
|  |  | Calculated LOD: |  | 0.3416 | mg/L | (calculated from the greater of LODs and LODb) |  |  |  |  | 2018-02-03 | 0.0001
|  |  | Calculated LOQ: |  | 1.1387 | mg/L | ( = 10/3 x LOD) |  |  |  |  | 2018-02-04 | 0.00221
[7 empty] | 0.02 | 0.08 | NO |  |  | 2018-02-05 | 0
16 | 0 | 0% | Existing LOD: |  | 0.04 | mg/L | 132 | 1 | 0.7576% | YES |  | 2018-02-06 | 0.0001
| Can the existing LOD be left unchanged? |  |  |  | NO |  | Use the new calculated LOD. |  |  |  |  | 2018-02-07 | 0.0048
|  |  |  | Is spike level okay (<5% spiked blanks <0)? | YES | [6 empty] | 2018-02-08 | 0.0067
[12 empty] | 2018-02-09 | 0.00181
Were outliers rejected? |  | Yes | (only gross failures may be excluded) | [8 empty] | 2018-02-10 | 0.0071
If so, explain: |  | Mislabeled vial from 10/4/19 - did not use in calculations | [9 empty] | 2018-02-11 | 0.003
[12 empty] | 2018-02-12 | 0.0012
NOTES: | [11 empty] | 2018-02-13 | 0.0011
Recalculate LODs and LODb at least every 13 months. | [11 empty] | 2018-02-14 | 0.0054
Include spiked blank data generated within the last 2 years as long as the data all used the same spike level (include initial results if within 2 years). ALL LINES DO NOT HAVE TO BE FILLED IN. | [11 empty] | 2018-02-15 | 0.001
[12 empty] | 2018-02-16 | 0.0015
If the lab thinks the sensitivity of the method has changed significantly, then the most recent data may be used (min. of 7 reps, 3 batches, over 3 days). | [11 empty] | 2018-02-17 | 0.0015
Include method blank data generated within the last 2 years. The lab has the option to use only the last 6 months or the 50 most recent method blanks, whichever yeilds the greater number of blanks. | [11 empty] | 2018-02-18 | 0.02
[12 empty] | 2018-02-19 | 0.001
The existing LOD can be left unchanged if the calculated LOD is within 0.5 - 2 times the existing LOD AND less than 3% of the method blanks are greater than the existing LOD. | [11 empty] | 2018-02-20 | 0.001
[12 empty] | 2018-02-21 | 0.0017
If sample analysis is not performed often for a given test, only quarters with actual sample analysis (do not include PT samples) need to have the spiked blanks analyzed. Take the last 24 months of data to generate the LOD. | [11 empty] | 2018-02-22 | 0.0007
[12 empty] | 2018-02-23 | 0.00998
If the method is altered in a way that could be expected to change its sensitivity, then re-determine the initial LOD and restart the ongoing data collection. | [11 empty] | 2018-02-24 | 0.00128
Only use data associated with passing calibrations and passing batch QC (reported data). | [11 empty] | 2018-02-25 | 0.0001
*The 2019 NR 149 code requires the LOQ be 10/3 the LOD or at the concentration of the lowest standard in the initial calibration. | [11 empty] | 2018-02-26 | 0.00221
[12 empty] | 2018-02-27 | 0
[12 empty] | 2018-02-28 | 0.0001
[12 empty] | 2018-03-01 | 0.0048
[12 empty] | 2018-03-02 | 0.0067
[12 empty] | 2018-03-03 | 0.00181
[12 empty] | 2018-03-04 | 0.0071
[12 empty] | 2018-03-05 | 0.003
[12 empty] | 2018-03-06 | 0.0012
[12 empty] | 2018-03-07 | 0.0011
[12 empty] | 2018-03-08 | 0.0054
[12 empty] | 2018-03-09 | 0.001
[12 empty] | 2018-03-10 | 0.0015
[12 empty] | 2018-03-11 | 0.0015
[12 empty] | 2018-03-12 | 0.02
[12 empty] | 2018-03-13 | 0.001
[12 empty] | 2018-03-14 | 0.001
[12 empty] | 2018-03-15 | 0.0017
[12 empty] | 2018-03-16 | 0.0007
[12 empty] | 2018-03-17 | 0.00998
[12 empty] | 2018-03-18 | 0.00128
[12 empty] | 2018-03-19 | 0.0001
[12 empty] | 2018-03-20 | 0.00221
[12 empty] | 2018-03-21 | 0
[12 empty] | 2018-03-22 | 0.0001
[12 empty] | 2018-03-23 | 0.0048
[12 empty] | 2018-03-24 | 0.0067
[12 empty] | 2018-03-25 | 0.00181
[12 empty] | 2018-03-26 | 0.0071
[12 empty] | 2018-03-27 | 0.003
[12 empty] | 2018-03-28 | 0.0012
[12 empty] | 2018-03-29 | 0.0011
[12 empty] | 2018-03-30 | 0.0054
[12 empty] | 2018-03-31 | 0.001
[12 empty] | 2018-04-01 | 0.0015
[12 empty] | 2018-04-02 | 0.0015
[12 empty] | 2018-04-03 | 0.02
[12 empty] | 2018-04-04 | 0.001
[12 empty] | 2018-04-05 | 0.001
[12 empty] | 2018-04-06 | 0.0017
[12 empty] | 2018-04-07 | 0.0007
[12 empty] | 2018-04-08 | 0.00998
[12 empty] | 2018-04-09 | 0.00128
[12 empty] | 2018-04-10 | 0.0001
[12 empty] | 2018-04-11 | 0.00221
[12 empty] | 2018-04-12 | 0
[12 empty] | 2018-04-13 | 0.0001
[12 empty] | 2018-04-14 | 0.0048
[12 empty] | 2018-04-15 | 0.0067
[12 empty] | 2018-04-16 | 0.00181
[12 empty] | 2018-04-17 | 0.0071
[12 empty] | 2018-04-18 | 0.003
[12 empty] | 2018-04-19 | 0.0012
[12 empty] | 2018-04-20 | 0.0011
[12 empty] | 2018-04-21 | 0.0054
[12 empty] | 2018-04-22 | 0.001
[12 empty] | 2018-04-23 | 0.0015
[12 empty] | 2018-04-24 | 0.0015
[12 empty] | 2018-04-25 | 0.02
[12 empty] | 2018-04-26 | 0.001
[12 empty] | 2018-04-27 | 0.001
[12 empty] | 2018-04-28 | 0.0017
[12 empty] | 2018-04-29 | 0.0007
[12 empty] | 2018-04-30 | 0.00998
[12 empty] | 2018-05-01 | 0.00128
[12 empty] | 2018-05-02 | 0.0001
[12 empty] | 2018-05-03 | 0.00221
[12 empty] | 2018-05-04 | 0
[12 empty] | 2018-05-05 | 0.0001
[12 empty] | 2018-05-06 | 0.0048
[12 empty] | 2018-05-07 | 0.0067
[12 empty] | 2018-05-08 | 0.00181
[12 empty] | 2018-05-09 | 0.0071
[12 empty] | 2018-05-10 | 0.003
*The 2019 NR 149 code requires the LOQ be 10/3 the LOD or at the concentration of the lowest standard in the initial calibration.
Fill in all of the white-colored spots as needed.
A spiked blank may be prepped and analyzed the same day.
For the initial LOD, there must be at least three different preparation days and at least three different analysis days.
Click on the cell and select either "99th percentile" (if more than 100 method blanks are entered) or "Standard Deviation."
It is up to the lab how they want to populate the method blank data (fill it in with each method blank, fill it in quarterly, upload from LIMS, etc.).
For the initial LODs only, the message, "repeat at a higher spike concentration," will appear if a result is negative.
Click on the cell and select either "Yes" or "No." If yes, you must explain why the data was rejected.
The print area is set up only to print one page (neither the notes nor the method blank data will be printed).
Only include the Inital LOD data if analyzed within the last 2 years.
The two spiked blanks per quarter must be prepared and analyzed in separate batches.
A message, "Use only the last 2 years of data," will appear if more than that is entered.
It is acceptable to have a "NO" in this cell--that just means the new on-going LOD must be used instead of the existing (old) LOD.
For chromatographic analyses, if a method blank was non-detect, enter "ND" instead of "0." LODb is determined differently if some or all method blanks are ND.

## Student's t-values
Number of Replicates | Degrees of Freedom (n - 1) | t(n-1, 0.99) | TINV(0.02,(A2-1))
7 | 6 | 3.143 | 3.143
8 | 7 | 2.998 | 2.998
9 | 8 | 2.896 | 2.896
10 | 9 | 2.821 | 2.821
11 | 10 | 2.764 | 2.764
12 | 11 | 2.718 | 2.718
13 | 12 | 2.681 | 2.681
14 | 13 | 2.65 | 2.65
15 | 14 | 2.624 | 2.624
16 | 15 | 2.602 | 2.602
17 | 16 | 2.583 | 2.583
18 | 17 | 2.567 | 2.567
19 | 18 | 2.552 | 2.552
20 | 19 | 2.539 | 2.539
21 | 20 | 2.528 | 2.528
22 | 21 | 2.518 | 2.518
23 | 22 | 2.508 | 2.508
24 | 23 | 2.5 | 2.5
25 | 24 | 2.492 | 2.492
26 | 25 | 2.485 | 2.485
27 | 26 | 2.479 | 2.479
28 | 27 | 2.473 | 2.473
29 | 28 | 2.467 | 2.467
30 | 29 | 2.462 | 2.462
31 | 30 | 2.457 | 2.457
32 | 31 | 2.453 | 2.453
33 | 32 | 2.449 | 2.449
34 | 33 | 2.445 | 2.445
35 | 34 | 2.441 | 2.441
36 | 35 | 2.438 | 2.438
37 | 36 | 2.434 | 2.434
38 | 37 | 2.431 | 2.431
39 | 38 | 2.429 | 2.429
40 | 39 | 2.426 | 2.426
41 | 40 | 2.423 | 2.423
42 | 41 | 2.421 | 2.421
43 | 42 | 2.418 | 2.418
44 | 43 | 2.416 | 2.416
45 | 44 | 2.414 | 2.414
46 | 45 | 2.412 | 2.412
47 | 46 | 2.41 | 2.41
48 | 47 | 2.408 | 2.408
49 | 48 | 2.407 | 2.407
50 | 49 | 2.405 | 2.405
51 | 50 | 2.403 | 2.403
52 | 51 | 2.402 | 2.402
53 | 52 | 2.4 | 2.4
54 | 53 | 2.399 | 2.399
55 | 54 | 2.397 | 2.397
56 | 55 | 2.396 | 2.396
57 | 56 | 2.395 | 2.395
58 | 57 | 2.394 | 2.394
59 | 58 | 2.392 | 2.392
60 | 59 | 2.391 | 2.391
61 | 60 | 2.39 | 2.39
62 | 61 | 2.389 | 2.389
63 | 62 | 2.388 | 2.388
64 | 63 | 2.387 | 2.387
65 | 64 | 2.386 | 2.386
66 | 65 | 2.385 | 2.385
67 | 66 | 2.384 | 2.384
68 | 67 | 2.383 | 2.383
69 | 68 | 2.382 | 2.382
70 | 69 | 2.382 | 2.382
71 | 70 | 2.381 | 2.381
72 | 71 | 2.38 | 2.38
73 | 72 | 2.379 | 2.379
74 | 73 | 2.379 | 2.379
75 | 74 | 2.378 | 2.378
76 | 75 | 2.377 | 2.377
77 | 76 | 2.376 | 2.376
78 | 77 | 2.376 | 2.376
79 | 78 | 2.375 | 2.375
80 | 79 | 2.374 | 2.374
81 | 80 | 2.374 | 2.374
82 | 81 | 2.373 | 2.373
83 | 82 | 2.373 | 2.373
84 | 83 | 2.372 | 2.372
85 | 84 | 2.372 | 2.372
86 | 85 | 2.371 | 2.371
87 | 86 | 2.37 | 2.37
88 | 87 | 2.37 | 2.37
89 | 88 | 2.369 | 2.369
90 | 89 | 2.369 | 2.369
91 | 90 | 2.368 | 2.368
92 | 91 | 2.368 | 2.368
93 | 92 | 2.368 | 2.368
94 | 93 | 2.367 | 2.367
95 | 94 | 2.367 | 2.367
96 | 95 | 2.366 | 2.366
97 | 96 | 2.366 | 2.366
98 | 97 | 2.365 | 2.365
99 | 98 | 2.365 | 2.365
100 | 99 | 2.365 | 2.365
101 | 100 | 2.3642 | 2.364
102 | 101 | 2.3638 | 2.364
103 | 102 | 2.3635 | 2.363
104 | 103 | 2.3631 | 2.363
105 | 104 | 2.3627 | 2.363
106 | 105 | 2.3624 | 2.362
107 | 106 | 2.362 | 2.362
108 | 107 | 2.3617 | 2.362
109 | 108 | 2.3614 | 2.361
110 | 109 | 2.3611 | 2.361
111 | 110 | 2.3607 | 2.361
112 | 111 | 2.3604 | 2.36
113 | 112 | 2.3601 | 2.36
114 | 113 | 2.3598 | 2.36
115 | 114 | 2.3595 | 2.36
116 | 115 | 2.3592 | 2.359
117 | 116 | 2.3589 | 2.359
118 | 117 | 2.3586 | 2.359
119 | 118 | 2.3583 | 2.358
120 | 119 | 2.3581 | 2.358
121 | 120 | 2.3578 | 2.358
122 | 121 | 2.3576 | 2.358
123 | 122 | 2.3573 | 2.357
124 | 123 | 2.3571 | 2.357
125 | 124 | 2.3568 | 2.357
126 | 125 | 2.3565 | 2.357
127 | 126 | 2.3563 | 2.356
128 | 127 | 2.3561 | 2.356
129 | 128 | 2.3559 | 2.356
130 | 129 | 2.3556 | 2.356
131 | 130 | 2.3554 | 2.355
132 | 131 | 2.3552 | 2.355
133 | 132 | 2.3549 | 2.355
134 | 133 | 2.3547 | 2.355
135 | 134 | 2.3545 | 2.354
136 | 135 | 2.3543 | 2.354
137 | 136 | 2.3541 | 2.354
138 | 137 | 2.3539 | 2.354
139 | 138 | 2.3537 | 2.354
140 | 139 | 2.3535 | 2.353
141 | 140 | 2.3533 | 2.353
142 | 141 | 2.3531 | 2.353
143 | 142 | 2.3529 | 2.353
144 | 143 | 2.3527 | 2.353
145 | 144 | 2.3525 | 2.353
146 | 145 | 2.3523 | 2.352
147 | 146 | 2.3522 | 2.352
148 | 147 | 2.352 | 2.352
149 | 148 | 2.3518 | 2.352
150 | 149 | 2.3516 | 2.352
151 | 150 | 2.3515 | 2.351
152 | 151 | 2.3513 | 2.351
153 | 152 | 2.3511 | 2.351
154 | 153 | 2.351 | 2.351
155 | 154 | 2.3508 | 2.351
156 | 155 | 2.3507 | 2.351
157 | 156 | 2.3505 | 2.35
158 | 157 | 2.3503 | 2.35
159 | 158 | 2.3502 | 2.35
160 | 159 | 2.35 | 2.35
161 | 160 | 2.3499 | 2.35
162 | 161 | 2.3497 | 2.35
163 | 162 | 2.3496 | 2.35
164 | 163 | 2.3495 | 2.349
165 | 164 | 2.3493 | 2.349
166 | 165 | 2.3492 | 2.349
167 | 166 | 2.349 | 2.349
168 | 167 | 2.3489 | 2.349
169 | 168 | 2.3487 | 2.349
170 | 169 | 2.3486 | 2.349
171 | 170 | 2.3485 | 2.348
172 | 171 | 2.3484 | 2.348
173 | 172 | 2.3482 | 2.348
174 | 173 | 2.3481 | 2.348
175 | 174 | 2.348 | 2.348
176 | 175 | 2.3478 | 2.348
177 | 176 | 2.3477 | 2.348
178 | 177 | 2.3476 | 2.348
179 | 178 | 2.3475 | 2.347
180 | 179 | 2.3474 | 2.347
181 | 180 | 2.3472 | 2.347
182 | 181 | 2.3471 | 2.347
183 | 182 | 2.347 | 2.347
184 | 183 | 2.3469 | 2.347
185 | 184 | 2.3468 | 2.347
186 | 185 | 2.3467 | 2.347
187 | 186 | 2.3466 | 2.347
188 | 187 | 2.3465 | 2.346
189 | 188 | 2.3463 | 2.346
190 | 189 | 2.3463 | 2.346
191 | 190 | 2.3461 | 2.346
192 | 191 | 2.346 | 2.346
193 | 192 | 2.3459 | 2.346
194 | 193 | 2.3458 | 2.346
195 | 194 | 2.3457 | 2.346
196 | 195 | 2.3456 | 2.346
197 | 196 | 2.3455 | 2.346
198 | 197 | 2.3454 | 2.345
199 | 198 | 2.3453 | 2.345
200 | 199 | 2.3452 | 2.345
201 | 200 | 2.3451 | 2.345
∞ | ∞ | 2.326

## Rev Record (hidden)
Revision Record
Date | Initials | Changes
2023-06-30 | ARF | Added tab for revision record. Added alert if more than 2 years of data is entered. Updated notes that not all lines need to be filled in and to use 2 years of method blank data. Removed footers and changed header to include info that this is only a guide. Editorial revisions.
2024-06-18 | ARF | Added alert if a value of 0 is entered for any method blank. Added a note to the NOTES tab to be sure to enter "ND" instead of 0 for chromatography. Added alert if more than 2 years of data is entered on the initial LOD tabs.