Lab Activities Frequency [Version 6/30/23]

| Lab Activities Reminder Table |  |
| --- | --- |
| One Time | IDC for each analyst and backup analyst for each test |
| 5 years | Weight must be sent out for external verification or replaced (class 2 or better) |
| Yearly | Thermometers must be replaced one year after the calibration date if the vendor has not provided a longer expiration date OR thermometers must be verified against a traceable thermometer |
|  | Temperature of the digestion block must be verified |
|  | PT sample for each test must be run and reported prior to August 31 each year |
|  | Barometer must be verified against an external source (monthly is best) |
|  | Perform LOD calculations and verifications |
|  | Recommendation: Check that all records are filed properly in labeled areas and folders |
|  | Recommendation: Review the SOPs and QM for corrections and that any changes needed are officially updated |
| Quarterly | Run two spiked blanks for the ongoing LODs (and could compile method blanks for the quarter too) |
|  | Mechanical pipettes must be verified |
|  | Check the pH of preserved samples for TP and NH3 (if there is a large industrial component to samples or if the plant processes have changed significantly, check pH more frequently) |
| Monthly | Balances must be checked with a weight and the weight must meet tolerances |
|  | Recommendation: File bench sheets and eDMR records in the month/year folder and put in labeled file cabinet |
| Daily | Record temperatures of refrigerators, ovens, and incubators |
| As it occurs | Corrective actions are documented |
|  | Maintenance is documented |
|  | Chemicals received and chemicals prepared are entered into the log; when new reagents are opened or disposed of, add dates |
|  | Each day the pH meter is used, calibrate the meter |
|  | Each day the DO meter is used, calibrate the meter |
|  | Each day the ammonia meter/probe is used, calibrate the meter |
|  | Calibrate the spectrophotometer after non-routine maintenance, CCVs fails twice, there is a change in its expected behavior, or if it leaves the direct control of the lab |

It is recommended that the laboratory use the Microsoft Outlook calendar reminder function or a wall calendar to ensure that these events are performed as required.

These are not all of the requirements that may need to be performed.  This form is only a guide, and it is the lab’s responsibility to ensure that requirements are met.