## Benchsheet
Nitrate Benchsheet (Colorimetry) [version 6/19/2024]
*** Fill in green shaded cells. ***
Laboratory:
| Analysis Method: |  |  | Hach 10206
| Analysis Date: |  |  | Analysis Time:
Analyst:
| Chemical Traceability | [7 empty] | Calibration Parameters
|  |  | Lot/ID # | Expiration Date |  | Concentration (mg/L) |  |  |  | Date of calibration curve:
| Hach Vials: | [11 empty] | Criteria
| Stock Standard (for curve): | [7 empty] | Correlation (r)= |  |  |  | ≥0.995
| Stock Standard for CCV: | [7 empty] | Slope=
| Second Source for ICV: | [7 empty] | y intercept= |  | x intercept=
| Matrix Spike Standard: | [7 empty] | LOD= |  | Date of LOD calc'n:
| Sample | pH adjust to 6-8? | Sample Vol. (mLs) | Sample + DI Vol. (mLs) | Dilution Factor | Absorbance | NO3-N (mg/L) | Final NO3-N (mg/L) |  | True Value (mg/L) |  | Quality Control
| Standard 0 (Cal. Blank) |  | 1 | 1 | 1
| Standard 1 |  | 1 | 1 | 1
| Standard 2 |  | 1 | 1 | 1
| Standard 3 |  | 1 | 1 | 1
| Standard 4 |  | 1 | 1 | 1
| Standard 5 |  | 1 | 1 | 1
| Standard 6 |  | 1 | 1 | 1
| Standard 7 |  | 1 | 1 | 1
| Standard 8 |  | 1 | 1 | 1
| ICV |  | 1 | 1 | 1
Criteria: 90-110%
| CCV 1 (low) |  | 1 | 1 | 1 |  |  |  |  | 0.3
Criteria: 85-115%
| Method Blank |  | 1 | 1 | 1 |  |  |  |  | Criteria: <LOD or <0.5 mg/L
| LCS/CCV 2 (mid/high) |  | 1 | 1 | 1 |  |  |  |  | 5
Criteria: 85-115%
| Matrix Spike | [8 empty] | Criteria: 85-115%
| Matrix Spike Duplicate | [8 empty] | Criteria: 20% RPD
Criteria:
| Spiked Sample ID: |  |  |  | #N/A | Lab Sample Result: |  | #N/A |  | Spike Volume Used (mL):
Final Volume (mL):
• The expiration dates will turn red if the chemical is expired. | [12 empty] | #DIV/0!
• The absorbance value entered will turn red if it is above the calibration range. The sample must then be diluted and re-run. | [12 empty] | #DIV/0!
• Hach Method 835 uses up to 1 mL of sample and the final volume must be 1 mL. Any difference between sample and final volume must be made up with reagent water.
• The ICV must be run immediately after generating the calibration curve. | [12 empty] | #DIV/0!
• On non-calibration days, the CCV must be run first, followed by the method blank.
• The order of analysis is from the top to bottom except the MS and MSD are analyzed immediately following the sample that was spiked.
• The method blank must be less than the highest of: limit of detection, 5% of the regulatory limit (10 mg/L), or 10% of sample concentration.
• If samples were acid preserved, must report result as nitrate+nitrite.
• The LOD calculation date will turn red as a warning if it is more than one year past the analysis date. Make sure LODs are calculated every 13 months.

## Standards Prep
Calibration Standards Preparation [version 6/19/24]
*** Fill in green shaded cells. ***
Concentration of the stock nitrate standard: |  |  | mg/L | mg/L = ppm = ug/mL
| Concentration of Standard to prepare (mg/L) | Final volume (mL) | Amount of stock standard to pipete (mL) | Amount of DI water to add (mL)
Standard 1
Standard 2
Standard 3
Standard 4
Standard 5
Standard 6
Standard 7
Standard 8
• You don't have to make 8 standards, but you have to make at least 3. At least 5 standards are recommended.
• Because the Hach vials used can range from 0.23 - 13.5 mg/L, standard 1 should be a concentration of about 0.2 mg/L and the highest standard must be no higher than 13.5 mg/L.
• Prepare the standards in volumetric flasks. Depending on what size of volumetric flasks you have and what size pipettes you have, you may have to play around to get pipette volumes that match what you have in your lab. You don't have to measure out the amount of DI water to add if you make up the solutions in volumetric flasks--just fill up to the line with DI water.

## IDC
Hach 10206 IDC [version 6/19/24]
Analyst: |  | Date of analysis:
| Result (mg/L) | Recovery (%)
5.0 mg/L IPR Std #1 |  | 0
5.0 mg/L IPR Std #2 |  | 0
5.0 mg/L IPR Std #3 |  | 0
5.0 mg/L IPR Std #4 |  | 0 | Limit
Average: | #DIV/0! | 0 | 90 - 110%
Standard Deviation: | #DIV/0! | 0
RSD: | #DIV/0! | #DIV/0! | <10%
Analyst: |  | Date of analysis:
| Result (mg/L) | Recovery (%)
5.0 mg/L IPR Std #1 |  | 0
5.0 mg/L IPR Std #2 |  | 0
5.0 mg/L IPR Std #3 |  | 0
5.0 mg/L IPR Std #4 |  | 0 | Limit
Average: | #DIV/0! | 0 | 90 - 110%
Standard Deviation: | #DIV/0! | 0
RSD: | #DIV/0! | #DIV/0! | <10%
Analyst: |  | Date of analysis:
| Result (mg/L) | Recovery (%)
5.0 mg/L IPR Std #1 |  | 0
5.0 mg/L IPR Std #2 |  | 0
5.0 mg/L IPR Std #3 |  | 0
5.0 mg/L IPR Std #4 |  | 0 | Limit
Average: | #DIV/0! | 0 | 90 - 110%
Standard Deviation: | #DIV/0! | 0
RSD: | #DIV/0! | #DIV/0! | <10%
• This spreadsheet is designed for using an IPR of 5 mg/L. If a lower concentration is used, contact Lab Cert.

## Rev Record (hidden)
Revision Record
Date | Initials | Changes
2023-06-30 | ARF | Added tab for revision record. Added revision dates to IDC and Stds Prep tabs. Fixed equations for alerting to enter in the LOD value. Added a column to record neutralization and added a comment at the bottom. Added x-intercept calculation and note if >LOD. Added LOD calculation date and warning. Editorial revisions.
2023-07-24 | ARF | Merged the "Spiked Sample ID" cell so only one cell has the drop down. Formulas were not affected.
2024-06-19 | ARF | Fixed the date format on IDC tab. Added equation to show the adjusted LOD for all dilutions. Removed part of note indicating neutralization was required for preserved samples.