## Quarterly Procedure
Quarterly Pipette Verification (gravimetric)* [Version 6/30/23]
Date |  | Analyst |  | Balance
Temp. (C) |  | Z-Factor | #N/A | Serial #
Pipette # |  | Pipette # |  | Pipette #
VOLUME | WEIGHT | VOLUME | WEIGHT | VOLUME | WEIGHT
(ml) | (g) | (ml) | (g) | (ml) | (g)
1 |  | 0.2 |  |  |  | 100 | 100 | #DIV/0!
1 |  | 0.2 |  |  |  | 100 | 100 | #DIV/0!
1 |  | 0.2 |  |  |  | 100 | 100 | #DIV/0!
1 |  | 0.2 |  |  |  | 100 | 100 | #DIV/0!
MEAN | #DIV/0! | MEAN | #DIV/0! | MEAN | #DIV/0!
CORR. MEAN | #DIV/0! | CORR. MEAN | #DIV/0! | CORR. MEAN | #DIV/0!
STD DEV | #DIV/0! | STD DEV | #DIV/0! | STD DEV | #DIV/0!
% CV | #DIV/0! | % CV | #DIV/0! | % CV | #DIV/0!
% INACC. | #DIV/0! | % INACC. | #DIV/0! | % INACC. | #DIV/0!
PASS / FAIL ? | #DIV/0! | PASS / FAIL ? | #DIV/0! | PASS / FAIL ? | #DIV/0!
Pipette # |  | Pipette # |  | Pipette #
VOLUME | WEIGHT | VOLUME | WEIGHT | VOLUME | WEIGHT
(ml) | (g) | (ml) | (g) | (ml) | (g)
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0!
MEAN | #DIV/0! | MEAN | #DIV/0! | MEAN | #DIV/0!
CORR. MEAN | #DIV/0! | CORR. MEAN | #DIV/0! | CORR. MEAN | #DIV/0!
STD DEV | #DIV/0! | STD DEV | #DIV/0! | STD DEV | #DIV/0!
% CV | #DIV/0! | % CV | #DIV/0! | % CV | #DIV/0!
% INACC. | #DIV/0! | % INACC. | #DIV/0! | % INACC. | #DIV/0!
PASS / FAIL ? | #DIV/0! | PASS / FAIL ? | #DIV/0! | PASS / FAIL ? | #DIV/0!
Corrective Action Due To Failures?
Acceptance Criteria
% Inaccuracy must be less than 2.00
No replicate may be greater than 2% from true volume.
%CV must be less than 1.00
*Created by the Wisconsin State Laboratory of Hygiene; WI DNR supplied form. This spreadsheet is only a guide and it is responsibility of the user to ensure that accurate results are reported.
15 | 1.002
15.5 | 1.002
16 | 1.0021
16.5 | 1.0022
17 | 1.0023
17.5 | 1.0024
18 | 1.0025
18.5 | 1.0026
19 | 1.0027
19.5 | 1.0028
20 | 1.0029
20.5 | 1.003
21 | 1.0031
21.5 | 1.0032
22 | 1.0033
22.5 | 1.0034
23 | 1.0035
23.5 | 1.0036
24 | 1.0037
24.5 | 1.0038
25 | 1.0039

## Instructions
1. Tare a beaker half filled with DI water.
a. Having water already in the beaker may help with any static electricity issues.
2. If using an adjustable pipette, set the pipette to the volume to be verified.
a. If using the pipette at several different volumes, it is recommended to verify the pipette at the low, mid-, and high range of the pipette.
b. If using the pipette at just one or two volumes, it is recommended to verify the pipette at the volumes that will actually be used for testing.
3. Pipette DI water into the beaker in the balance.
4. Record the weight, then tare the balance again.
5. Repeat steps 3 and 4 for a total of four measurements.
6. Enter the values into the spreadsheet. The spreadsheet will determine if the pipette meets all criteria.
a. If the pipette fails, refer to the pipette manual for maintenance and/or recalibration.
b. After taking the corrective action, repeat the four measurements. If the pipette fails again, do not use for compliance samples.

## Lookup Tables
Temperature
ºC | Z-Factor
15 | 1.002
15.5 | 1.002
16 | 1.0021
16.5 | 1.0022
17 | 1.0023
17.5 | 1.0024
18 | 1.0025
18.5 | 1.0026
19 | 1.0027
19.5 | 1.0028
20 | 1.0029
20.5 | 1.003
21 | 1.0031
21.5 | 1.0032
22 | 1.0033
22.5 | 1.0034
23 | 1.0035
23.5 | 1.0036
24 | 1.0037
24.5 | 1.0038
25 | 1.0039

## Rev Record (hidden)
Revision Record
Date | Initials | Changes
2023-06-30 | ARF | Added tab for revision record.