## Benchsheet
Total Phosphorus Benchsheet (Colorimetry) [Version 6/30/23]
*** Fill in green shaded cells. ***
Laboratory: |  |  |  |  | Digestion Method: |  |  |  |  | Analysis Method:
Phosphorus Permit Limit:
Analysis Traceability
Digestion Date |  |  |  | Analysis Date
Digestion Analyst |  |  | Analyst
Digestion Block Temp (°C) |  |  | (Check the temperature each use)
Autoclave Pressure (kPa) |  |  | (Check the pressure each use)
Chemical Traceability
|  | Concentration | Vendor |  | Lot # | Expiration Date
|  | (mg/L as P) | [5 empty] | Calibration Parameters
Stock Standard | [8 empty] | Date of calibration curve:
Working Standard/CCV | [11 empty] | Criteria
Second Source Std (ICV) | [7 empty] | Correlation (r)= |  |  |  | ≥0.995
Potassium Antimonyl Tartrate | [7 empty] | Slope=
Ammonium Molybdate | [7 empty] | y Intercept= |  | x intercept=
Ascorbic Acid | [7 empty] | LOD=
Hach TNT Vials | [7 empty] | Date of LOD Calculation:
Sample | pH adjust 6-8 (Y/N)? | Sample Vol. (mLs) | Sample + DI Vol. (mLs) | Dilution Factor | Absorbance | Total Phosphorus (mg/L) | Final Total Phosphorus (mg/L) |  | True Value (mg/L) |  | Quality Control
Calibration Blank |  |  |  | 1 |  |  |  |  | 0
Standard 1 |  |  |  | 1
Standard 2 |  |  |  | 1
Standard 3 |  |  |  | 1
Standard 4 |  |  |  | 1
Standard 5 |  |  |  | 1
Standard 6 |  |  |  | 1
Standard 7 |  |  |  | 1
Standard 8 |  |  |  | 1
ICV
Criteria: 90-110%
LCS/CCV
Criteria: 90-110%
Method Blank | [8 empty] | Criteria: <LOD
CCV
Criteria: 90-110%
• The expiration dates will turn red if the chemical is expired.
• The absorbance value entered will turn red if it is above the calibration range. The sample must then be diluted and re-run.
• The LOD calculation date will turn red as a warning if it is more than one year past the analysis date. Make sure LODs are calculated every 13 months.

## Standards Prep
Calibration Standards Preparation REVISED 6/30/23
Concentration of the stock phosporus standard: |  | 10 | mg/L | mg/L = ppm = ug/mL
| Concentration of Standard to prepare (mg/L) | Final volume (mL) | Amount of stock standard to pipete (mL) | Amount of DI water to add (mL)
Standard 1 | 0.2 | 5 | 0.1 | 4.9
Standard 2 | 0.4 | 5 | 0.2 | 4.8
Standard 3 | 0.6 | 5 | 0.3 | 4.7
Standard 4 | 0.8 | 5 | 0.4 | 4.6
Standard 5 | 1 | 5 | 0.5 | 4.5
Standard 6
Standard 7
Standard 8

## Rev Record (hidden)
Revision Record
Date | Initials | Changes
2023-06-30 | ARF | Added tab for revision record. Added revision dates to Stds Prep tabs. Added note for LOD calculation date and warning. Removed note and formating for expired calibration curve. Editorial revisions.